ACSM Best Practices

This document contains recommendations for setup, operation, maintenance and data analysis procedures for the Aerodyne aerosol chemical speciation monitor (ACSM). It is intended as a supplement to the published ACSM DAQ and Igor manuals. Information in this reference is drawn from the ACSM DAQ and Igor manuals, the Aerodyne ACSM website (https://sites.google.com/site/ariacsm/home), the ToF-AMS wiki page (http://cires.colorado.edu/jimenez-group/wiki/index.php?title=ToF-AMS_Main), ACSM user experiences, and discussion during the May 2012 ACSM meeting at PSI.

1. ACSM Resources

ACSM DAQ and Igor manual

Aerodyne’s ACSM website

https://sites.google.com/site/ariacsm/home

ToF-AMS wiki page

http://cires.colorado.edu/jimenez-group/wiki/index.php?title=ToF-AMS_Main

Site intended for ToF-AMS, but many issues are similar, e.g. collection efficiency, sampling line considerations, and instrument inter-comparison issues, and some operating/data analysis principles.

ACTRIS campaign website

http://www.psi.ch/acsm-stations/acsm-and-emep-stations

2. Inlet and Sampling Lines

Construction of sampling system

The complete sampling system should be diagrammed, including sampling line lengths, diameters, materials, and flow rates. Please put the information on the ACTRIS-PSI-Disk

Sample flow controller accessory (Aerodyne)

Several of the issues below (cyclone, pump/flow rate measurement) are facilitated by the sample flow controller accessory available from Aerodyne.

PM2.5 Cyclone on sampling  inlet

Reduces clogging of the ACSM critical orifice.  PM2.5 is preferred to PM1 because supermicron particles may become submicron when dried as recommended.

Construct sampling lines from stainless and/or copper tubing

Conductive tubing materials reduce particle losses.  Critically, black conductive tubing (e.g from TSI) should not be used, because this can cause siloxane contamination of the mass spectrum, including at m/z 73 (an important marker for biomass burning).  Stainless steel is ideal for long-term sampling.  If copper is used, lines should be periodically checked for corrosion.

Line losses and leaks

Residence time in the sampling lines should be limited to reduce diffusive particle losses. This can be accomplished with (1) a pump (e.g. split from the Nafion dryer pump below) and/or (2) connecting other instrumentation (e.g. SMPS) near the ACSM inlet.  Required flowrate depends on sampling lines, generally 1 to 10 L/min up to the dryer inlet.  This flowrate should be continuously monitored and recorded.

The combination of tube width and sampling flow should be selected to maintain laminar flow (Re < 1000) to avoid particle losses from turbulent flow.

Important to test sampling line for leaks (e.g. filter on or near sampling inlet).

Nafion dryer

ACSM bounce collection efficiency is affected by relative humidity in ways that are difficult to predict. Note that it is the RH at the ACSM inlet that matters, not the ambient RH, and thus this affect is also influenced by outdoor/indoor temperature gradients. It is strongly recommended to dry the sampled aerosol (25 to 50% RH) before the ACSM inlet. This allows the collection efficiency to be estimated from the particle composition, and also facilitates comparisons with co-located instrumentation (e.g. SMPS).

Nafion dryers are preferred over silica gel dryers because they require less maintenance (silica gel dryers may require daily gel changes). The dryer should have a metal casing to reduce charged particle losses. The PermaPure MD series dryers work well:

http://www.permapure.com/products/dryers/md-gas-dryers/

The aerosol should be plumbed through the inner path, and a pump with critical orifice connected to the outer path to reduce the outer path pressure to ~0.1 atm. When connecting the inner path, care should be taken with the fittings to avoid breaking the internal seals on the nafion membrane, introducing a leak to atmosphere. This can be checked by placing a filter at the dryer inlet and measuring zero particles at the dryer outlet with a CPC.

More information on dryer setup and operation is available at:

http://cires.colorado.edu/jimenez-group/wiki/index.php/Field_ToF-AMS_Operation#Dryer_Set_Up

2. Spare parts and support equipment

On-line uninterruptible power supply (UPS, 2000 kV)

Servo motor for inlet assembly

Filter for valve body assembly.  Filter should be changed when a pressure drop is observed across the filter (see diagnostics plot in ACSM analysis software).

SEM (monitor the required voltage as a function of time to anticipate failure).

Filaments (2 sets of 2).  At any time, 2 filaments are installed in the instrument.  Should compare their performance after installation.

Critical orifice

Turbo pump

Diaphragms for backing pump and tool for changing diaphragms

3. Maintenance/Calibration Schedule

Continuously monitor instrument flowrate, signals for m/z 28 and naphthalene, and turbo pump speeds/currents/temperatures.  Stability of these parameters generally indicates stable ACSM operation.  However, it is still advisable to regularly perform calibration/maintenance activities, particularly for a new instrument, at the beginning of a new deployment, or following venting of the instrument.  For an instrument for which long-term stability has been confirmed, the calibration schedule can be reduced.

2x per week

Monitor inlet pressure (~flowrate), m/z 28 and naphthalene signals.  Significant changes indicate maintenance is needed.

Monitor turbo pump speeds, currents, and temperatures.  This helps to anticipate pump failure.

Check and record SEM gain, adjust if needed.

Monitor Faraday signal.

For SEM and Faraday, daily checks are best at the start of a campaign or after venting the instrument, until stable operation is observed.

Every month (or as needed; this is highly site-dependent)

Clean dust from turbo pump fan filters.

Check flow rate, clean orifice with isopropanol, check flow rate again.

Every 2 months

IE calibration with ammonium nitrate.  This should be done more frequently (weekly) at the start of a campaign or after the instrument has been vented, until stable results (i.e airbeam to IE ratio) are obtained.

Every 3 months

Collect data with (1) valve closed to check background (2) valve open with filter in-line.  Several data points are needed, suggest collecting for at least 1-2 hrs at 15-min time resolution.

Check inlet pressure with valve open and closed to evaluate vacuum/pumping integrity.

Keep DAQ software updated (~6 months or following crucial update).  DAQ is available from Aerodyne’s ACSM users’ site.

Other

Evaluate stability of m/z calibration and resolution.  Normally this is stable, but should be individually confirmed.

4. Calibrations

Flow calibration

Critical to record the ambient pressure and temperature at which flow calibration is performed.

Necessary to adjust critical orifice size if instrument is operated at altitude.

Ionization efficiency

All IE calibrations should be immediately preceded by SEM gain measurement.

It is important that the NH4NO3 calibration solution not be more concentrated than recommended (5mM) to avoid artifacts from multiply-charged particles.

Confirm that particles are dried before size-selection.

Measure ions for NO3+ (m/z 30, 46), NH4+ (15, 16, 17), naphthalene (m/z=128), and air (32 and 40, but NOT 28).  The m/z 28 signal should be instead taken from the SEM gain measurement.  Measuring m/z 28 during the IE calibration shortens the SEM lifetime.

Ratio of m/z 28 to IE should be logged.  This serves as a reference point for sensitivity (SEM gain, MS tuning…) corrections during data analysis.

Naphthalene signal should be recorded as a function of SEM gain.  The ACSM temperature must also be logged to obtain a temperature-corrected naphthalene signal.

Relative ionization efficiency

SO42-: Should be measured immediately following an IE calibration, where the RIE of NH4+ has just been calculated from mass balance between NH4+ and NO3-.  Nebulize (NH4)2SO4 and sample with the ACSM.  The RIE of SO42- can then be calculated from mass balance with NH4+, since NH4+’s RIE is known.  This method is not biased by particle bounce.

Methanesulfonic acid (MSA): It is advisable to calibrate for MSA if sampling near the ocean.  This species has been observed in previous ACSM/AMS measurements but the RIE is not well known.

5. Operation and analysis

Sampling time

30 minute time resolution is generally recommended.  Maximum of 1 hr for sites with very low concentrations.  15 minutes can be very helpful if local sources are important.  Note that short sampling times can easily be re-gridded during post-analysis to generate longer averages if signal-to-noise is problematic.

Collection efficiency (from particle bounce)

Uncertainties due to bounce collection efficiency (Eb) can be greatly reduced by installing a Nafion dryer as discussed above.

Two methods of estimating the bounce collection efficiency are available: (1) theoretical estimate based using ACSM-measured particle composition; and (2) comparison with external instruments (e.g. SMPS). Both should be investigated.

Details of the composition-based estimate (based on acidity and nitrate content) are given in Middlebrook et al. (2012). The paper uses AMS data, but applies equally to the ACSM:

Middlebrook, A.M., Bahreini, R., Jimenez, J.L., and Canagaratna, M.R.: Evaluation of composition-dependent collection efficiencies for the Aerodyne aerosol mass spectrometer using field data, Aerosol Sci. Technol., 46, 258-271, doi:10.1080/02786826.2011.620041, 2012.

http://cires.colorado.edu/jimenez/Papers/2011_AST_Middlebrook_CE.pdf

The collection efficiency can also be assessed through comparison with co-located instruments, such as the SMPS. Care should be taken to assess assumptions/approximations used to perform such inter-comparisons! Note also that other instruments have their own uncertainties, and disagreements with the ACSM may not necessarily be due to collection efficiency.  Collection efficiency should not be a catch-all for disagreements between instruments!

A few points relating to comparisons of ACSM with SMPS (or similar instruments):

(1) Particles should be dried for both instruments

(2) Transmission ranges of the ACSM and SMPS should be synchronized, especially at large particle sizes (can be done during SMPS analysis, as long as the large-size SMPS cutpoint exceeds that of the ACSM)

(3) Refractory species (black carbon, dust, sea salt, etc.) must be negligible or measured by other methods with comparable size transmission.

Sensitivity/gain corrections during analysis

Recommended to use m/z 28 signal (rather than naphthalene) for gain/sensitivity corrections due to better signal-to-noise.